## summaryO3 18 Jun 16
Averages of total ozone:
With the exisiting Calibration, Lamp tests included
Average is over the complete Mu range
D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD
270.5014 | 286.0086 | 270.7227 | 258.5586 | 266.838
3.6634 | 19.1706 | 3.8847 | -8.2794
With Langley Plot Calibration
Langley Plot Adjustments and Lamp Test results included
D083_TO3_C* | D083_TO3_D* | D083_TO3_A* | D083_TO3_CD* | D083_TO3_AD*
264.2779 | 258.0345 | 264.3672 | 269.166 | 265.9472
-1.6693 | -7.9127 | -1.58 | 3.2188

## IdD083
Dobson 083 at IZO 2016.Sep.20 Langley Plot Results
SLC083Na | -0.9 | SLC083Nc | -0.8 | [9 empty] | SLC083Nd | -0.9 | [11 empty] | Correction to
D083 LP adj. A:-1.77 C:-0.98 D:-1.57 Nad:-0.20 | [12 empty] | MaxMu | 2.5 | MaxMu can be changed | [5 empty] | LPC_C_83 |  | LPC_D_83 |  | LPC_A_83 | AD | CD
[20 empty] | With out High mu values | -1.1105 |  | -2.146 |  | -2.1745 | -0.028578 | 1.0355 | [14 empty] | Limited Mu range
[11 empty] | 46 |  |  |  | 270.5014 | 286.0086 | 270.7227 | 258.5586 | 266.838 | all Values | -0.97595 |  | -1.5666 |  | -1.7678 | -0.20123 | 0.59064 | [14 empty] | 1.0286 |  | -0.019364
D083RC | D083RD | D083RA | day | h | m | s | m | year |  |  |  | D083NC | D083ND | D083NA | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | 1/mu.C | P*C_83 | 1/Mu.D | P*D_83 | 1/Mu.A | P*A_83 | lst.C | Mu.C | AirM.C | lst.D | Mu.D | AirM.D | lst.A | Mu.A | AirM.A | lst.CD | Mu.CD | AirM.CD | lst.AD | Mu.AD | AirM.AD | 1/mu.cd | P*CD_83 | 1/mu.ad | P*AD_83
146.6 | 84.5 | 258.6 | 20 | 7 | 52 | 0 | 9 | 16 | C | 83 | IZO | 137.2144 | 80.2644 | 245.7916 | 266.8137 | 276.3656 | 267.0687 | 259.7557 | 264.6102 | 0.22438 | 0.22052 | 0.22782 | 0.09966 | 0.23127 | 0.47741 | 2016-09-20 07:52 | 4.4567 | 4.7003 | 2016-09-20 07:53 | 4.3894 | 4.6209 | 2016-09-20 07:54 | 4.3239 | 4.5441 | 2016-09-20 07:52 | 4.4228 | 4.6602 | 2016-09-20 07:54 | 4.3564 | 4.5821 | 0.2261 | 0.12086 | 0.22955 | 0.37775
139.3 | 79.4 | 245.3 | 20 | 7 | 56 | 0 | 9 | 16 | C | 83 | IZO | 130.3337 | 75.578 | 232.3154 | 270.6149 | 277.2089 | 267.9297 | 265.7873 | 265.4943 | 0.23819 | 0.22366 | 0.24165 | 0.099965 | 0.24511 | 0.47895 | 2016-09-20 07:56 | 4.1984 | 4.3981 | 2016-09-20 07:57 | 4.1383 | 4.3286 | 2016-09-20 07:58 | 4.0798 | 4.2613 | 2016-09-20 07:56 | 4.1681 | 4.363 | 2016-09-20 07:57 | 4.1088 | 4.2947 | 0.23992 | 0.1237 | 0.24338 | 0.37898
130.3 | 74.8 | 233.5 | 20 | 8 | 0 | 0 | 9 | 16 | C | 83 | IZO | 121.8649 | 71.3628 | 220.1665 | 267.2392 | 277.6474 | 268.6313 | 259.4134 | 266.2768 | 0.25205 | 0.22087 | 0.25552 | 0.10012 | 0.259 | 0.4802 | 2016-09-20 08:00 | 3.9675 | 4.1329 | 2016-09-20 08:01 | 3.9136 | 4.0716 | 2016-09-20 08:02 | 3.8611 | 4.0122 | 2016-09-20 08:01 | 3.9403 | 4.102 | 2016-09-20 08:02 | 3.8871 | 4.0417 | 0.25379 | 0.12075 | 0.25726 | 0.38008
123.2 | 70.7 | 222.6 | 20 | 8 | 4 | 0 | 9 | 16 | C | 83 | IZO | 115.2438 | 67.6156 | 209.1018 | 266.9181 | 278.1927 | 269.0877 | 258.3854 | 266.7221 | 0.26595 | 0.22061 | 0.26943 | 0.10032 | 0.27291 | 0.48102 | 2016-09-20 08:04 | 3.7601 | 3.8985 | 2016-09-20 08:05 | 3.7116 | 3.8441 | 2016-09-20 08:06 | 3.6643 | 3.7913 | 2016-09-20 08:04 | 3.7357 | 3.8711 | 2016-09-20 08:06 | 3.6878 | 3.8175 | 0.26769 | 0.12029 | 0.27117 | 0.3807
117 | 67.1 | 212.2 | 20 | 8 | 8 | 0 | 9 | 16 | C | 83 | IZO | 109.482 | 64.3347 | 198.7114 | 267.2219 | 279.3707 | 268.8764 | 257.9834 | 266.1689 | 0.27986 | 0.22086 | 0.28334 | 0.10074 | 0.28682 | 0.48064 | 2016-09-20 08:08 | 3.5732 | 3.69 | 2016-09-20 08:09 | 3.5293 | 3.6414 | 2016-09-20 08:10 | 3.4865 | 3.5941 | 2016-09-20 08:08 | 3.5511 | 3.6655 | 2016-09-20 08:10 | 3.5078 | 3.6176 | 0.2816 | 0.12012 | 0.28508 | 0.37989
111.6 | 63.6 | 202.8 | 20 | 8 | 12 | 0 | 9 | 16 | C | 83 | IZO | 104.4817 | 61.1298 | 189.4174 | 268.2267 | 278.503 | 268.8536 | 260.4142 | 266.3671 | 0.29377 | 0.22169 | 0.29725 | 0.10043 | 0.30072 | 0.4806 | 2016-09-20 08:12 | 3.404 | 3.5033 | 2016-09-20 08:13 | 3.3642 | 3.4597 | 2016-09-20 08:14 | 3.3253 | 3.4172 | 2016-09-20 08:12 | 3.384 | 3.4814 | 2016-09-20 08:13 | 3.3446 | 3.4383 | 0.29551 | 0.12126 | 0.29899 | 0.38017
106.6 | 60.8 | 194.5 | 20 | 8 | 16 | 0 | 9 | 16 | C | 83 | IZO | 99.8483 | 58.5542 | 181.2463 | 268.8418 | 280.4535 | 269.3145 | 259.9766 | 266.4576 | 0.30767 | 0.2222 | 0.31114 | 0.10113 | 0.3146 | 0.48142 | 2016-09-20 08:16 | 3.2503 | 3.3354 | 2016-09-20 08:17 | 3.214 | 3.296 | 2016-09-20 08:18 | 3.1786 | 3.2576 | 2016-09-20 08:17 | 3.232 | 3.3156 | 2016-09-20 08:18 | 3.1962 | 3.2767 | 0.3094 | 0.12106 | 0.31287 | 0.38029
101.7 | 58.2 | 187 | 20 | 8 | 20 | 0 | 9 | 16 | C | 83 | IZO | 95.2972 | 56.15 | 173.8576 | 268.1428 | 281.8261 | 269.8897 | 257.6583 | 266.8363 | 0.32153 | 0.22162 | 0.32499 | 0.10163 | 0.32844 | 0.48245 | 2016-09-20 08:20 | 3.1101 | 3.1835 | 2016-09-20 08:21 | 3.077 | 3.1478 | 2016-09-20 08:22 | 3.0447 | 3.113 | 2016-09-20 08:20 | 3.0935 | 3.1656 | 2016-09-20 08:22 | 3.0608 | 3.1303 | 0.32326 | 0.11999 | 0.32672 | 0.38082
97.6 | 55.6 | 179.6 | 20 | 8 | 24 | 0 | 9 | 16 | C | 83 | IZO | 91.4858 | 53.7352 | 166.5579 | 268.8371 | 281.1138 | 269.4312 | 259.4282 | 266.4461 | 0.33535 | 0.22219 | 0.33879 | 0.10137 | 0.34224 | 0.48163 | 2016-09-20 08:24 | 2.982 | 3.0456 | 2016-09-20 08:25 | 2.9516 | 3.0131 | 2016-09-20 08:26 | 2.922 | 2.9814 | 2016-09-20 08:24 | 2.9667 | 3.0293 | 2016-09-20 08:26 | 2.9367 | 2.9972 | 0.33707 | 0.12082 | 0.34052 | 0.38026
93.7 | 53.1 | 173 | 20 | 8 | 28 | 0 | 9 | 16 | C | 83 | IZO | 87.8711 | 51.3978 | 160.1077 | 269.0026 | 279.169 | 269.4732 | 261.2162 | 266.9938 | 0.34911 | 0.22233 | 0.35254 | 0.10067 | 0.35597 | 0.4817 | 2016-09-20 08:28 | 2.8644 | 2.9199 | 2016-09-20 08:29 | 2.8365 | 2.8902 | 2016-09-20 08:30 | 2.8092 | 2.8611 | 2016-09-20 08:28 | 2.8504 | 2.905 | 2016-09-20 08:29 | 2.8228 | 2.8756 | 0.35083 | 0.12166 | 0.35426 | 0.38103
90 | 51.2 | 166.8 | 20 | 8 | 32 | 0 | 9 | 16 | C | 83 | IZO | 84.4419 | 49.6071 | 154.0459 | 268.6903 | 280.8566 | 269.2588 | 259.3466 | 266.3017 | 0.36281 | 0.22207 | 0.36623 | 0.10128 | 0.36964 | 0.48132 | 2016-09-20 08:32 | 2.7562 | 2.8049 | 2016-09-20 08:33 | 2.7305 | 2.7776 | 2016-09-20 08:34 | 2.7053 | 2.7509 | 2016-09-20 08:33 | 2.7433 | 2.7912 | 2016-09-20 08:34 | 2.7179 | 2.7642 | 0.36452 | 0.12079 | 0.36793 | 0.38004
86.7 | 49.2 | 161.6 | 20 | 8 | 36 | 0 | 9 | 16 | C | 83 | IZO | 81.3815 | 47.7108 | 148.9574 | 268.8306 | 279.9315 | 270.1325 | 260.3037 | 267.6324 | 0.37644 | 0.22219 | 0.37984 | 0.10095 | 0.38323 | 0.48288 | 2016-09-20 08:36 | 2.6564 | 2.6992 | 2016-09-20 08:37 | 2.6327 | 2.6742 | 2016-09-20 08:38 | 2.6094 | 2.6496 | 2016-09-20 08:36 | 2.6445 | 2.6866 | 2016-09-20 08:38 | 2.621 | 2.6618 | 0.37814 | 0.12124 | 0.38154 | 0.38194
83.6 | 47.5 | 156.2 | 20 | 8 | 40 | 0 | 9 | 16 | C | 83 | IZO | 78.5181 | 46.089 | 143.7213 | 268.7984 | 280.4307 | 269.8297 | 259.8534 | 267.1291 | 0.39 | 0.22216 | 0.39337 | 0.10113 | 0.39674 | 0.48234 | 2016-09-20 08:40 | 2.5641 | 2.602 | 2016-09-20 08:41 | 2.5421 | 2.5788 | 2016-09-20 08:42 | 2.5205 | 2.5562 | 2016-09-20 08:40 | 2.5531 | 2.5903 | 2016-09-20 08:42 | 2.5313 | 2.5675 | 0.39168 | 0.12104 | 0.39506 | 0.38122
80.7 | 46 | 151.3 | 20 | 8 | 44 | 0 | 9 | 16 | C | 83 | IZO | 75.8504 | 44.65 | 138.9706 | 268.688 | 281.6342 | 269.7727 | 258.7203 | 266.7555 | 0.40346 | 0.22207 | 0.40681 | 0.10156 | 0.41016 | 0.48224 | 2016-09-20 08:44 | 2.4786 | 2.5121 | 2016-09-20 08:45 | 2.4581 | 2.4907 | 2016-09-20 08:46 | 2.4381 | 2.4698 | 2016-09-20 08:44 | 2.4683 | 2.5014 | 2016-09-20 08:46 | 2.4481 | 2.4802 | 0.40514 | 0.12051 | 0.40848 | 0.38068
78.2 | 44.5 | 146.6 | 20 | 8 | 48 | 0 | 9 | 16 | C | 83 | IZO | 73.55 | 43.2039 | 134.4389 | 269.465 | 281.5402 | 269.4424 | 260.1664 | 266.3671 | 0.41683 | 0.22271 | 0.42016 | 0.10153 | 0.42348 | 0.48165 | 2016-09-20 08:48 | 2.3991 | 2.429 | 2016-09-20 08:49 | 2.3801 | 2.4091 | 2016-09-20 08:50 | 2.3614 | 2.3896 | 2016-09-20 08:49 | 2.3895 | 2.419 | 2016-09-20 08:50 | 2.3707 | 2.3993 | 0.41849 | 0.12119 | 0.42182 | 0.38012
75.7 | 43.2 | 142.6 | 20 | 8 | 52 | 0 | 9 | 16 | C | 83 | IZO | 71.2466 | 41.9447 | 130.617 | 269.3846 | 282.4818 | 270.0904 | 259.2905 | 266.9422 | 0.4301 | 0.22265 | 0.4334 | 0.10187 | 0.43669 | 0.48281 | 2016-09-20 08:52 | 2.325 | 2.3517 | 2016-09-20 08:53 | 2.3073 | 2.3333 | 2016-09-20 08:54 | 2.2899 | 2.3152 | 2016-09-20 08:52 | 2.3162 | 2.3425 | 2016-09-20 08:54 | 2.2986 | 2.3242 | 0.43175 | 0.12078 | 0.43505 | 0.38094
73.5 | 41.9 | 138.6 | 20 | 8 | 56 | 0 | 9 | 16 | C | 83 | IZO | 69.2191 | 40.6799 | 126.8028 | 269.9454 | 282.2851 | 270.1105 | 260.4339 | 267.0184 | 0.44326 | 0.22311 | 0.44653 | 0.1018 | 0.4498 | 0.48284 | 2016-09-20 08:56 | 2.256 | 2.2799 | 2016-09-20 08:57 | 2.2395 | 2.2627 | 2016-09-20 08:58 | 2.2232 | 2.2459 | 2016-09-20 08:56 | 2.2477 | 2.2713 | 2016-09-20 08:58 | 2.2313 | 2.2543 | 0.4449 | 0.12132 | 0.44817 | 0.38105
69.3 | 39.6 | 131.4 | 20 | 9 | 4 | 0 | 9 | 16 | C | 83 | IZO | 65.3509 | 38.4269 | 119.9283 | 269.8952 | 282.3342 | 270.2327 | 260.3005 | 267.1611 | 0.46924 | 0.22307 | 0.47246 | 0.10181 | 0.47566 | 0.48306 | 2016-09-20 09:04 | 2.1311 | 2.1504 | 2016-09-20 09:05 | 2.1166 | 2.1355 | 2016-09-20 09:06 | 2.1023 | 2.1207 | 2016-09-20 09:05 | 2.1238 | 2.1429 | 2016-09-20 09:06 | 2.1094 | 2.1281 | 0.47085 | 0.12126 | 0.47406 | 0.38125
67.3 | 38.5 | 128.2 | 20 | 9 | 8 | 0 | 9 | 16 | C | 83 | IZO | 63.5093 | 37.3419 | 116.8749 | 269.3465 | 281.49 | 270.4705 | 259.978 | 267.6732 | 0.48205 | 0.22262 | 0.48523 | 0.10151 | 0.4884 | 0.48349 | 2016-09-20 09:08 | 2.0745 | 2.0919 | 2016-09-20 09:09 | 2.0609 | 2.0779 | 2016-09-20 09:10 | 2.0475 | 2.0641 | 2016-09-20 09:08 | 2.0677 | 2.0849 | 2016-09-20 09:10 | 2.0542 | 2.071 | 0.48364 | 0.12111 | 0.48682 | 0.38198
65.7 | 37.6 | 125.1 | 20 | 9 | 12 | 0 | 9 | 16 | C | 83 | IZO | 62.0334 | 36.4508 | 113.9228 | 270.3149 | 282.4097 | 270.47 | 260.9818 | 267.4409 | 0.49473 | 0.22342 | 0.49788 | 0.10184 | 0.50102 | 0.48349 | 2016-09-20 09:12 | 2.0213 | 2.0371 | 2016-09-20 09:13 | 2.0085 | 2.024 | 2016-09-20 09:14 | 1.9959 | 2.011 | 2016-09-20 09:12 | 2.0149 | 2.0305 | 2016-09-20 09:14 | 2.0022 | 2.0174 | 0.4963 | 0.12158 | 0.49945 | 0.38165
63.9 | 36.6 | 122.1 | 20 | 9 | 16 | 0 | 9 | 16 | C | 83 | IZO | 60.3697 | 35.4573 | 111.0838 | 269.5737 | 281.1124 | 270.2909 | 260.6689 | 267.545 | 0.50727 | 0.2228 | 0.51039 | 0.10137 | 0.51349 | 0.48317 | 2016-09-20 09:16 | 1.9713 | 1.9857 | 2016-09-20 09:17 | 1.9593 | 1.9733 | 2016-09-20 09:18 | 1.9475 | 1.9611 | 2016-09-20 09:16 | 1.9653 | 1.9794 | 2016-09-20 09:18 | 1.9534 | 1.9672 | 0.50883 | 0.12143 | 0.51194 | 0.38179
62.6 | 35.9 | 119.3 | 20 | 9 | 20 | 0 | 9 | 16 | C | 83 | IZO | 59.1667 | 34.7601 | 108.4539 | 271.1216 | 283.2756 | 270.2431 | 261.7387 | 266.9391 | 0.51968 | 0.22408 | 0.52276 | 0.10215 | 0.52583 | 0.48308 | 2016-09-20 09:20 | 1.9243 | 1.9373 | 2016-09-20 09:21 | 1.9129 | 1.9256 | 2016-09-20 09:22 | 1.9018 | 1.9142 | 2016-09-20 09:20 | 1.9186 | 1.9314 | 2016-09-20 09:22 | 1.9073 | 1.9199 | 0.52122 | 0.12193 | 0.52429 | 0.38093
61.2 | 35.1 | 116.7 | 20 | 9 | 24 | 0 | 9 | 16 | C | 83 | IZO | 57.8693 | 33.9615 | 106.0145 | 271.596 | 283.3121 | 270.3173 | 262.5504 | 267.0234 | 0.53194 | 0.22447 | 0.53498 | 0.10217 | 0.53801 | 0.48321 | 2016-09-20 09:24 | 1.8799 | 1.8918 | 2016-09-20 09:25 | 1.8692 | 1.8808 | 2016-09-20 09:26 | 1.8587 | 1.87 | 2016-09-20 09:24 | 1.8746 | 1.8863 | 2016-09-20 09:26 | 1.8639 | 1.8754 | 0.53346 | 0.12231 | 0.5365 | 0.38105
59.6 | 34.4 | 114.2 | 20 | 9 | 28 | 0 | 9 | 16 | C | 83 | IZO | 56.3835 | 33.2614 | 103.6642 | 270.3393 | 284.1563 | 270.2441 | 259.6661 | 266.7189 | 0.54405 | 0.22344 | 0.54706 | 0.10247 | 0.55005 | 0.48308 | 2016-09-20 09:28 | 1.8381 | 1.8489 | 2016-09-20 09:29 | 1.828 | 1.8385 | 2016-09-20 09:30 | 1.818 | 1.8283 | 2016-09-20 09:28 | 1.833 | 1.8437 | 2016-09-20 09:30 | 1.823 | 1.8334 | 0.54555 | 0.12097 | 0.54855 | 0.38061
58.3 | 33.5 | 111.9 | 20 | 9 | 32 | 0 | 9 | 16 | C | 83 | IZO | 55.1733 | 32.3597 | 101.5039 | 270.3973 | 281.2603 | 270.3623 | 262.0093 | 267.5992 | 0.55601 | 0.22348 | 0.55898 | 0.10143 | 0.56193 | 0.48329 | 2016-09-20 09:32 | 1.7985 | 1.8084 | 2016-09-20 09:33 | 1.789 | 1.7986 | 2016-09-20 09:34 | 1.7796 | 1.789 | 2016-09-20 09:32 | 1.7937 | 1.8035 | 2016-09-20 09:34 | 1.7843 | 1.7938 | 0.5575 | 0.12206 | 0.56046 | 0.38187
57.1 | 32.9 | 109.4 | 20 | 9 | 36 | 0 | 9 | 16 | C | 83 | IZO | 54.0551 | 31.7576 | 99.1658 | 270.631 | 282.3655 | 269.5327 | 261.5671 | 266.281 | 0.56782 | 0.22368 | 0.57074 | 0.10182 | 0.57366 | 0.48181 | 2016-09-20 09:36 | 1.7611 | 1.7701 | 2016-09-20 09:37 | 1.7521 | 1.7609 | 2016-09-20 09:38 | 1.7432 | 1.7518 | 2016-09-20 09:36 | 1.7566 | 1.7655 | 2016-09-20 09:38 | 1.7476 | 1.7563 | 0.56928 | 0.12185 | 0.5722 | 0.37999
56.1 | 32.4 | 107.5 | 20 | 9 | 40 | 0 | 9 | 16 | C | 83 | IZO | 53.1219 | 31.2552 | 97.3979 | 271.7368 | 284.5419 | 270.1737 | 261.8432 | 266.5344 | 0.57946 | 0.22459 | 0.58234 | 0.10261 | 0.58522 | 0.48296 | 2016-09-20 09:40 | 1.7258 | 1.734 | 2016-09-20 09:41 | 1.7172 | 1.7253 | 2016-09-20 09:42 | 1.7088 | 1.7166 | 2016-09-20 09:40 | 1.7215 | 1.7296 | 2016-09-20 09:42 | 1.713 | 1.7209 | 0.5809 | 0.12198 | 0.58378 | 0.38035
54.9 | 31.8 | 105.6 | 20 | 9 | 44 | 0 | 9 | 16 | C | 83 | IZO | 51.9998 | 30.6515 | 95.6309 | 271.1218 | 284.5797 | 270.4983 | 260.722 | 266.9318 | 0.59093 | 0.22408 | 0.59378 | 0.10262 | 0.59661 | 0.48354 | 2016-09-20 09:44 | 1.6922 | 1.6998 | 2016-09-20 09:45 | 1.6841 | 1.6915 | 2016-09-20 09:46 | 1.6761 | 1.6834 | 2016-09-20 09:44 | 1.6882 | 1.6956 | 2016-09-20 09:46 | 1.6801 | 1.6874 | 0.59236 | 0.12146 | 0.59519 | 0.38091
53.9 | 31.3 | 103.8 | 20 | 9 | 48 | 0 | 9 | 16 | C | 83 | IZO | 51.0622 | 30.1477 | 93.9509 | 271.4329 | 285.7635 | 270.8027 | 260.3566 | 267.0142 | 0.60224 | 0.22434 | 0.60504 | 0.10305 | 0.60783 | 0.48408 | 2016-09-20 09:48 | 1.6605 | 1.6674 | 2016-09-20 09:49 | 1.6528 | 1.6595 | 2016-09-20 09:50 | 1.6452 | 1.6518 | 2016-09-20 09:48 | 1.6566 | 1.6635 | 2016-09-20 09:50 | 1.649 | 1.6557 | 0.60364 | 0.12129 | 0.60644 | 0.38103
47.3 | 27.7 | 92.1 | 20 | 10 | 20 | 0 | 9 | 16 | C | 83 | IZO | 44.8083 | 26.4963 | 82.9492 | 271.5268 | 286.2486 | 272.0598 | 260.1428 | 268.4681 | 0.68615 | 0.22442 | 0.68857 | 0.10322 | 0.69098 | 0.48633 | 2016-09-20 10:20 | 1.4574 | 1.461 | 2016-09-20 10:21 | 1.4523 | 1.4558 | 2016-09-20 10:22 | 1.4472 | 1.4506 | 2016-09-20 10:20 | 1.4548 | 1.4583 | 2016-09-20 10:22 | 1.4497 | 1.4532 | 0.68736 | 0.12119 | 0.68978 | 0.3831
46.7 | 27.7 | 90.9 | 20 | 10 | 24 | 0 | 9 | 16 | C | 83 | IZO | 44.2328 | 26.4963 | 81.8257 | 271.9113 | 293.3115 | 272.0798 | 255.3575 | 266.7084 | 0.69577 | 0.22474 | 0.69814 | 0.10577 | 0.7005 | 0.48636 | 2016-09-20 10:24 | 1.4373 | 1.4405 | 2016-09-20 10:25 | 1.4324 | 1.4356 | 2016-09-20 10:26 | 1.4276 | 1.4307 | 2016-09-20 10:24 | 1.4348 | 1.4381 | 2016-09-20 10:25 | 1.43 | 1.4331 | 0.69695 | 0.11896 | 0.69932 | 0.38059
46.1 | 27.3 | 89.8 | 20 | 10 | 28 | 0 | 9 | 16 | C | 83 | IZO | 43.6555 | 26.0877 | 80.7993 | 272.0386 | 292.205 | 272.272 | 256.4393 | 267.2289 | 0.70518 | 0.22484 | 0.7075 | 0.10537 | 0.70981 | 0.48671 | 2016-09-20 10:28 | 1.4181 | 1.4211 | 2016-09-20 10:29 | 1.4134 | 1.4164 | 2016-09-20 10:30 | 1.4088 | 1.4117 | 2016-09-20 10:28 | 1.4157 | 1.4187 | 2016-09-20 10:30 | 1.4111 | 1.414 | 0.70634 | 0.11947 | 0.70865 | 0.38134
45.5 | 27.1 | 88.7 | 20 | 10 | 32 | 0 | 9 | 16 | C | 83 | IZO | 43.077 | 25.8832 | 79.7753 | 271.9138 | 294.7917 | 272.2737 | 254.2153 | 266.5775 | 0.71438 | 0.22474 | 0.71665 | 0.10631 | 0.71891 | 0.48671 | 2016-09-20 10:32 | 1.3998 | 1.4026 | 2016-09-20 10:33 | 1.3954 | 1.3981 | 2016-09-20 10:34 | 1.391 | 1.3937 | 2016-09-20 10:32 | 1.3976 | 1.4003 | 2016-09-20 10:34 | 1.3932 | 1.3959 | 0.71552 | 0.11843 | 0.71778 | 0.38041
44.9 | 26.7 | 87.8 | 20 | 10 | 36 | 0 | 9 | 16 | C | 83 | IZO | 42.4973 | 25.4738 | 78.9391 | 271.5405 | 292.9986 | 272.8412 | 254.9407 | 267.7417 | 0.72338 | 0.22443 | 0.72559 | 0.10566 | 0.72779 | 0.48773 | 2016-09-20 10:36 | 1.3824 | 1.385 | 2016-09-20 10:37 | 1.3782 | 1.3807 | 2016-09-20 10:38 | 1.374 | 1.3765 | 2016-09-20 10:36 | 1.3803 | 1.3828 | 2016-09-20 10:38 | 1.3761 | 1.3786 | 0.72448 | 0.11877 | 0.72669 | 0.38207
44.6 | 26.5 | 86.8 | 20 | 10 | 40 | 0 | 9 | 16 | C | 83 | IZO | 42.207 | 25.2688 | 78.0112 | 273.4954 | 295.0162 | 272.8536 | 256.8467 | 267.2474 | 0.73215 | 0.22604 | 0.73432 | 0.10639 | 0.73646 | 0.48775 | 2016-09-20 10:40 | 1.3658 | 1.3682 | 2016-09-20 10:41 | 1.3618 | 1.3641 | 2016-09-20 10:42 | 1.3578 | 1.3601 | 2016-09-20 10:40 | 1.3638 | 1.3662 | 2016-09-20 10:41 | 1.3598 | 1.3621 | 0.73324 | 0.11966 | 0.73539 | 0.38136
44.1 | 26.6 | 86.1 | 20 | 10 | 44 | 0 | 9 | 16 | C | 83 | IZO | 41.7224 | 25.3713 | 77.3622 | 273.537 | 303.1147 | 273.8441 | 250.6518 | 266.4417 | 0.74072 | 0.22608 | 0.74282 | 0.10931 | 0.74492 | 0.48952 | 2016-09-20 10:44 | 1.35 | 1.3522 | 2016-09-20 10:45 | 1.3462 | 1.3483 | 2016-09-20 10:46 | 1.3424 | 1.3445 | 2016-09-20 10:44 | 1.3481 | 1.3503 | 2016-09-20 10:46 | 1.3443 | 1.3464 | 0.74177 | 0.11677 | 0.74387 | 0.38021
43.7 | 26.2 | 85.1 | 20 | 10 | 48 | 0 | 9 | 16 | C | 83 | IZO | 41.334 | 24.9612 | 76.4357 | 274.2404 | 300.4272 | 273.5086 | 253.9795 | 266.7008 | 0.74906 | 0.22666 | 0.75111 | 0.10834 | 0.75315 | 0.48892 | 2016-09-20 10:48 | 1.335 | 1.337 | 2016-09-20 10:49 | 1.3314 | 1.3333 | 2016-09-20 10:50 | 1.3278 | 1.3297 | 2016-09-20 10:48 | 1.3332 | 1.3352 | 2016-09-20 10:50 | 1.3296 | 1.3315 | 0.75009 | 0.11832 | 0.75213 | 0.38058
43.2 | 26 | 84.3 | 20 | 10 | 52 | 0 | 9 | 16 | C | 83 | IZO | 40.8477 | 24.7559 | 75.695 | 273.8573 | 301.7112 | 273.7845 | 252.3058 | 266.7219 | 0.75718 | 0.22634 | 0.75918 | 0.1088 | 0.76116 | 0.48941 | 2016-09-20 10:52 | 1.3207 | 1.3225 | 2016-09-20 10:53 | 1.3172 | 1.319 | 2016-09-20 10:54 | 1.3138 | 1.3156 | 2016-09-20 10:52 | 1.3189 | 1.3208 | 2016-09-20 10:54 | 1.3155 | 1.3173 | 0.75818 | 0.11754 | 0.76017 | 0.38061
42.7 | 25.7 | 83.4 | 20 | 10 | 56 | 0 | 9 | 16 | C | 83 | IZO | 40.3604 | 24.4477 | 74.8633 | 273.259 | 300.5601 | 273.5078 | 252.1351 | 266.6663 | 0.76508 | 0.22585 | 0.76702 | 0.10839 | 0.76894 | 0.48892 | 2016-09-20 10:56 | 1.3071 | 1.3088 | 2016-09-20 10:57 | 1.3037 | 1.3054 | 2016-09-20 10:58 | 1.3005 | 1.3021 | 2016-09-20 10:56 | 1.3054 | 1.3071 | 2016-09-20 10:57 | 1.3021 | 1.3038 | 0.76605 | 0.11746 | 0.76798 | 0.38053
42.3 | 25.4 | 82.5 | 20 | 11 | 0 | 0 | 9 | 16 | C | 83 | IZO | 39.9698 | 24.1393 | 74.033 | 273.3617 | 299.1149 | 273.0696 | 253.4357 | 266.4826 | 0.77275 | 0.22593 | 0.77463 | 0.10786 | 0.7765 | 0.48813 | 2016-09-20 11:00 | 1.2941 | 1.2956 | 2016-09-20 11:01 | 1.2909 | 1.2925 | 2016-09-20 11:02 | 1.2878 | 1.2893 | 2016-09-20 11:01 | 1.2925 | 1.2941 | 2016-09-20 11:02 | 1.2894 | 1.2909 | 0.77369 | 0.11807 | 0.77557 | 0.38027
41.9 | 25.1 | 81.9 | 20 | 11 | 4 | 0 | 9 | 16 | C | 83 | IZO | 39.5785 | 23.8305 | 73.4799 | 273.275 | 297.3798 | 273.682 | 254.6246 | 267.6883 | 0.78019 | 0.22586 | 0.78202 | 0.10724 | 0.78383 | 0.48923 | 2016-09-20 11:04 | 1.2817 | 1.2832 | 2016-09-20 11:05 | 1.2787 | 1.2802 | 2016-09-20 11:06 | 1.2758 | 1.2772 | 2016-09-20 11:04 | 1.2802 | 1.2817 | 2016-09-20 11:06 | 1.2773 | 1.2787 | 0.78111 | 0.11862 | 0.78292 | 0.38199
41.6 | 25 | 81.2 | 20 | 11 | 8 | 0 | 9 | 16 | C | 83 | IZO | 39.2845 | 23.7275 | 72.8348 | 273.9348 | 299.8687 | 273.7465 | 253.8685 | 267.1403 | 0.7874 | 0.22641 | 0.78917 | 0.10814 | 0.79092 | 0.48934 | 2016-09-20 11:08 | 1.27 | 1.2713 | 2016-09-20 11:09 | 1.2672 | 1.2685 | 2016-09-20 11:10 | 1.2644 | 1.2656 | 2016-09-20 11:08 | 1.2686 | 1.2699 | 2016-09-20 11:10 | 1.2658 | 1.267 | 0.78828 | 0.11827 | 0.79004 | 0.38121
41.5 | 24.8 | 80.6 | 20 | 11 | 12 | 0 | 9 | 16 | C | 83 | IZO | 39.1865 | 23.5215 | 72.2817 | 276.3149 | 299.8863 | 274.0756 | 258.077 | 267.5481 | 0.79437 | 0.22837 | 0.79608 | 0.10814 | 0.79777 | 0.48993 | 2016-09-20 11:12 | 1.2589 | 1.2601 | 2016-09-20 11:13 | 1.2562 | 1.2573 | 2016-09-20 11:14 | 1.2535 | 1.2547 | 2016-09-20 11:12 | 1.2575 | 1.2587 | 2016-09-20 11:13 | 1.2548 | 1.256 | 0.79523 | 0.12023 | 0.79693 | 0.38179
41 | 24.8 | 80 | 20 | 11 | 16 | 0 | 9 | 16 | C | 83 | IZO | 38.6955 | 23.5215 | 71.7284 | 274.7573 | 304.2578 | 274.2655 | 251.9299 | 266.6814 | 0.80111 | 0.22709 | 0.80276 | 0.10972 | 0.80439 | 0.49027 | 2016-09-20 11:16 | 1.2483 | 1.2494 | 2016-09-20 11:17 | 1.2457 | 1.2468 | 2016-09-20 11:18 | 1.2432 | 1.2442 | 2016-09-20 11:17 | 1.247 | 1.2481 | 2016-09-20 11:18 | 1.2444 | 1.2455 | 0.80194 | 0.11737 | 0.80358 | 0.38055
40.7 | 24.6 | 79.4 | 20 | 11 | 20 | 0 | 9 | 16 | C | 83 | IZO | 38.4003 | 23.3153 | 71.1747 | 274.9226 | 303.8455 | 274.3168 | 252.5421 | 266.8498 | 0.80761 | 0.22722 | 0.8092 | 0.10957 | 0.81077 | 0.49036 | 2016-09-20 11:20 | 1.2382 | 1.2392 | 2016-09-20 11:21 | 1.2358 | 1.2368 | 2016-09-20 11:22 | 1.2334 | 1.2344 | 2016-09-20 11:20 | 1.237 | 1.238 | 2016-09-20 11:22 | 1.2346 | 1.2356 | 0.80841 | 0.11765 | 0.80999 | 0.38079
40.4 | 24.4 | 78.9 | 20 | 11 | 24 | 0 | 9 | 16 | C | 83 | IZO | 38.1047 | 23.109 | 70.7128 | 274.9258 | 303.2013 | 274.6524 | 253.0463 | 267.4332 | 0.81387 | 0.22723 | 0.8154 | 0.10934 | 0.81691 | 0.49096 | 2016-09-20 11:24 | 1.2287 | 1.2296 | 2016-09-20 11:25 | 1.2264 | 1.2273 | 2016-09-20 11:26 | 1.2241 | 1.225 | 2016-09-20 11:24 | 1.2275 | 1.2285 | 2016-09-20 11:26 | 1.2253 | 1.2262 | 0.81464 | 0.11789 | 0.81615 | 0.38163

## IdD083_star
This Page uses the lamp tests and the langley plot results in the limited mu range from sheet IdD083 | [21 empty] | the correction values from ID083 are applied to the NC
Dobson 083 at IZO 2016.Sep.20 Langley Plot Results | [15 empty] | 109
SLC083Na | -0.9 | SLC083Nc | -0.7 | [9 empty] | SLC083Nd | -0.7
D083 LP adj. A:0.52 C:0.12 D:0.51 Nad:0.01 | [20 empty] | LPC_C_83 |  | LPC_D_83 |  | LPC_A_83
[15 empty] | 264.2779 | 258.0345 | 264.3672 | 269.166 | 265.9472 | All Obs | 0.1216 |  | 0.50576 |  | 0.52038 | 0.014621 | -0.38415
D083RC | D083RD | D083RA | day | h | m | s | m | year |  |  |  | D083NC | D083ND | D083NA | D083_TO3_C* | D083_TO3_D* | D083_TO3_A* | D083_TO3_CD* | D083_TO3_AD* | 1/mu.C | P*C_83 | 1/Mu.D | P*D_83 | 1/Mu.A | P*A_83 | lst.C | Mu.C | AirM.C | lst.D | Mu.D | AirM.D | lst.A | Mu.A | AirM.A | lst.CD | Mu.CD | AirM.CD | lst.AD | Mu.AD | AirM.AD
146.6 | 84.5 | 258.6 | 20 | 7 | 52 | 0 | 9 | 16 | C | 83 | IZO | 136.2039 | 78.3184 | 243.617 | 264.0704 | 264.0715 | 264.2553 | 264.405 | 264.1928 | 0.22438 | 0.21825 | 0.22782 | 0.095227 | 0.23127 | 0.47238 | 2016-09-20 07:52 | 4.4567 | 4.7003 | 2016-09-20 07:53 | 4.3894 | 4.6209 | 2016-09-20 07:54 | 4.3239 | 4.5441 | 2016-09-20 07:52 | 4.4228 | 4.6602 | 2016-09-20 07:54 | 4.3564 | 4.5821
139.3 | 79.4 | 245.3 | 20 | 7 | 56 | 0 | 9 | 16 | C | 83 | IZO | 129.3232 | 73.632 | 230.1408 | 267.7028 | 264.1689 | 264.948 | 270.7146 | 265.0544 | 0.23819 | 0.22126 | 0.24165 | 0.095262 | 0.24511 | 0.47362 | 2016-09-20 07:56 | 4.1984 | 4.3981 | 2016-09-20 07:57 | 4.1383 | 4.3286 | 2016-09-20 07:58 | 4.0798 | 4.2613 | 2016-09-20 07:56 | 4.1681 | 4.363 | 2016-09-20 07:57 | 4.1088 | 4.2947
130.3 | 74.8 | 233.5 | 20 | 8 | 0 | 0 | 9 | 16 | C | 83 | IZO | 120.8545 | 69.4169 | 217.9919 | 264.1576 | 263.8586 | 265.4807 | 264.6195 | 265.8146 | 0.25205 | 0.21833 | 0.25552 | 0.09515 | 0.259 | 0.47457 | 2016-09-20 08:00 | 3.9675 | 4.1329 | 2016-09-20 08:01 | 3.9136 | 4.0716 | 2016-09-20 08:02 | 3.8611 | 4.0122 | 2016-09-20 08:01 | 3.9403 | 4.102 | 2016-09-20 08:02 | 3.8871 | 4.0417
123.2 | 70.7 | 222.6 | 20 | 8 | 4 | 0 | 9 | 16 | C | 83 | IZO | 114.2333 | 65.6696 | 206.9273 | 263.6666 | 263.6536 | 265.7678 | 263.8707 | 266.2375 | 0.26595 | 0.21792 | 0.26943 | 0.095076 | 0.27291 | 0.47508 | 2016-09-20 08:04 | 3.7601 | 3.8985 | 2016-09-20 08:05 | 3.7116 | 3.8441 | 2016-09-20 08:06 | 3.6643 | 3.7913 | 2016-09-20 08:04 | 3.7357 | 3.8711 | 2016-09-20 08:06 | 3.6878 | 3.8175
117 | 67.1 | 212.2 | 20 | 8 | 8 | 0 | 9 | 16 | C | 83 | IZO | 108.4715 | 62.3888 | 196.5368 | 263.8002 | 264.0807 | 265.3873 | 263.7482 | 265.662 | 0.27986 | 0.21803 | 0.28334 | 0.09523 | 0.28682 | 0.4744 | 2016-09-20 08:08 | 3.5732 | 3.69 | 2016-09-20 08:09 | 3.5293 | 3.6414 | 2016-09-20 08:10 | 3.4865 | 3.5941 | 2016-09-20 08:08 | 3.5511 | 3.6655 | 2016-09-20 08:10 | 3.5078 | 3.6176
111.6 | 63.6 | 202.8 | 20 | 8 | 12 | 0 | 9 | 16 | C | 83 | IZO | 103.4712 | 59.1839 | 187.2428 | 264.6349 | 262.4625 | 265.1954 | 266.4582 | 265.838 | 0.29377 | 0.21872 | 0.29725 | 0.094647 | 0.30072 | 0.47406 | 2016-09-20 08:12 | 3.404 | 3.5033 | 2016-09-20 08:13 | 3.3642 | 3.4597 | 2016-09-20 08:14 | 3.3253 | 3.4172 | 2016-09-20 08:12 | 3.384 | 3.4814 | 2016-09-20 08:13 | 3.3446 | 3.4383
106.6 | 60.8 | 194.5 | 20 | 8 | 16 | 0 | 9 | 16 | C | 83 | IZO | 98.8378 | 56.6083 | 179.0718 | 265.0802 | 263.6637 | 265.4875 | 266.2993 | 265.9064 | 0.30767 | 0.21909 | 0.31114 | 0.09508 | 0.3146 | 0.47458 | 2016-09-20 08:16 | 3.2503 | 3.3354 | 2016-09-20 08:17 | 3.214 | 3.296 | 2016-09-20 08:18 | 3.1786 | 3.2576 | 2016-09-20 08:17 | 3.232 | 3.3156 | 2016-09-20 08:18 | 3.1962 | 3.2767
101.7 | 58.2 | 187 | 20 | 8 | 20 | 0 | 9 | 16 | C | 83 | IZO | 94.2867 | 54.204 | 171.6831 | 264.2118 | 264.2888 | 265.8943 | 264.2589 | 266.2631 | 0.32153 | 0.21837 | 0.32499 | 0.095305 | 0.32844 | 0.47531 | 2016-09-20 08:20 | 3.1101 | 3.1835 | 2016-09-20 08:21 | 3.077 | 3.1478 | 2016-09-20 08:22 | 3.0447 | 3.113 | 2016-09-20 08:20 | 3.0935 | 3.1656 | 2016-09-20 08:22 | 3.0608 | 3.1303
97.6 | 55.6 | 179.6 | 20 | 8 | 24 | 0 | 9 | 16 | C | 83 | IZO | 90.4753 | 51.7893 | 164.3833 | 264.737 | 262.8314 | 265.2679 | 266.3057 | 265.8509 | 0.33535 | 0.21881 | 0.33879 | 0.09478 | 0.34224 | 0.47419 | 2016-09-20 08:24 | 2.982 | 3.0456 | 2016-09-20 08:25 | 2.9516 | 3.0131 | 2016-09-20 08:26 | 2.922 | 2.9814 | 2016-09-20 08:24 | 2.9667 | 3.0293 | 2016-09-20 08:26 | 2.9367 | 2.9972
93.7 | 53.1 | 173 | 20 | 8 | 28 | 0 | 9 | 16 | C | 83 | IZO | 86.8606 | 49.4518 | 157.9331 | 264.7343 | 260.1446 | 265.1429 | 268.3695 | 266.3768 | 0.34911 | 0.2188 | 0.35254 | 0.093811 | 0.35597 | 0.47396 | 2016-09-20 08:28 | 2.8644 | 2.9199 | 2016-09-20 08:29 | 2.8365 | 2.8902 | 2016-09-20 08:30 | 2.8092 | 2.8611 | 2016-09-20 08:28 | 2.8504 | 2.905 | 2016-09-20 08:29 | 2.8228 | 2.8756
90 | 51.2 | 166.8 | 20 | 8 | 32 | 0 | 9 | 16 | C | 83 | IZO | 83.4314 | 47.6612 | 151.8714 | 264.2545 | 261.0938 | 264.7623 | 266.7743 | 265.6631 | 0.36281 | 0.21841 | 0.36623 | 0.094153 | 0.36964 | 0.47328 | 2016-09-20 08:32 | 2.7562 | 2.8049 | 2016-09-20 08:33 | 2.7305 | 2.7776 | 2016-09-20 08:34 | 2.7053 | 2.7509 | 2016-09-20 08:33 | 2.7433 | 2.7912 | 2016-09-20 08:34 | 2.7179 | 2.7642
86.7 | 49.2 | 161.6 | 20 | 8 | 36 | 0 | 9 | 16 | C | 83 | IZO | 80.371 | 45.7648 | 146.7829 | 264.2281 | 259.4342 | 265.4706 | 268.0043 | 266.9723 | 0.37644 | 0.21839 | 0.37984 | 0.093555 | 0.38323 | 0.47455 | 2016-09-20 08:36 | 2.6564 | 2.6992 | 2016-09-20 08:37 | 2.6327 | 2.6742 | 2016-09-20 08:38 | 2.6094 | 2.6496 | 2016-09-20 08:36 | 2.6445 | 2.6866 | 2016-09-20 08:38 | 2.621 | 2.6618
83.6 | 47.5 | 156.2 | 20 | 8 | 40 | 0 | 9 | 16 | C | 83 | IZO | 77.5076 | 44.1431 | 141.5467 | 264.0302 | 259.2032 | 265.0034 | 267.8252 | 266.4477 | 0.39 | 0.21822 | 0.39337 | 0.093471 | 0.39674 | 0.47371 | 2016-09-20 08:40 | 2.5641 | 2.602 | 2016-09-20 08:41 | 2.5421 | 2.5788 | 2016-09-20 08:42 | 2.5205 | 2.5562 | 2016-09-20 08:40 | 2.5531 | 2.5903 | 2016-09-20 08:42 | 2.5313 | 2.5675
80.7 | 46 | 151.3 | 20 | 8 | 44 | 0 | 9 | 16 | C | 83 | IZO | 74.8399 | 42.7041 | 136.7961 | 263.7553 | 259.6814 | 264.7832 | 266.9615 | 266.0528 | 0.40346 | 0.21799 | 0.40681 | 0.093644 | 0.41016 | 0.47332 | 2016-09-20 08:44 | 2.4786 | 2.5121 | 2016-09-20 08:45 | 2.4581 | 2.4907 | 2016-09-20 08:46 | 2.4381 | 2.4698 | 2016-09-20 08:44 | 2.4683 | 2.5014 | 2016-09-20 08:46 | 2.4481 | 2.4802
78.2 | 44.5 | 146.6 | 20 | 8 | 48 | 0 | 9 | 16 | C | 83 | IZO | 72.5395 | 41.2579 | 132.2643 | 264.3688 | 258.8673 | 264.2909 | 268.675 | 265.6435 | 0.41683 | 0.2185 | 0.42016 | 0.09335 | 0.42348 | 0.47244 | 2016-09-20 08:48 | 2.3991 | 2.429 | 2016-09-20 08:49 | 2.3801 | 2.4091 | 2016-09-20 08:50 | 2.3614 | 2.3896 | 2016-09-20 08:49 | 2.3895 | 2.419 | 2016-09-20 08:50 | 2.3707 | 2.3993
75.7 | 43.2 | 142.6 | 20 | 8 | 52 | 0 | 9 | 16 | C | 83 | IZO | 70.2361 | 39.9988 | 128.4424 | 264.1261 | 259.0943 | 264.7781 | 268.0645 | 266.1978 | 0.4301 | 0.2183 | 0.4334 | 0.093432 | 0.43669 | 0.47331 | 2016-09-20 08:52 | 2.325 | 2.3517 | 2016-09-20 08:53 | 2.3073 | 2.3333 | 2016-09-20 08:54 | 2.2899 | 2.3152 | 2016-09-20 08:52 | 2.3162 | 2.3425 | 2016-09-20 08:54 | 2.2986 | 2.3242
73.5 | 41.9 | 138.6 | 20 | 8 | 56 | 0 | 9 | 16 | C | 83 | IZO | 68.2086 | 38.7339 | 124.6283 | 264.5261 | 258.1888 | 264.6388 | 269.4711 | 266.2533 | 0.44326 | 0.21863 | 0.44653 | 0.093106 | 0.4498 | 0.47306 | 2016-09-20 08:56 | 2.256 | 2.2799 | 2016-09-20 08:57 | 2.2395 | 2.2627 | 2016-09-20 08:58 | 2.2232 | 2.2459 | 2016-09-20 08:56 | 2.2477 | 2.2713 | 2016-09-20 08:58 | 2.2313 | 2.2543
69.3 | 39.6 | 131.4 | 20 | 9 | 4 | 0 | 9 | 16 | C | 83 | IZO | 64.3404 | 36.4809 | 117.7537 | 264.1582 | 256.8391 | 264.4464 | 269.8569 | 266.3555 | 0.46924 | 0.21833 | 0.47246 | 0.092619 | 0.47566 | 0.47272 | 2016-09-20 09:04 | 2.1311 | 2.1504 | 2016-09-20 09:05 | 2.1166 | 2.1355 | 2016-09-20 09:06 | 2.1023 | 2.1207 | 2016-09-20 09:05 | 2.1238 | 2.1429 | 2016-09-20 09:06 | 2.1094 | 2.1281
67.3 | 38.5 | 128.2 | 20 | 9 | 8 | 0 | 9 | 16 | C | 83 | IZO | 62.4988 | 35.3959 | 114.7004 | 263.4529 | 255.3054 | 264.5292 | 269.7902 | 266.8476 | 0.48205 | 0.21774 | 0.48523 | 0.092066 | 0.4884 | 0.47287 | 2016-09-20 09:08 | 2.0745 | 2.0919 | 2016-09-20 09:09 | 2.0609 | 2.0779 | 2016-09-20 09:10 | 2.0475 | 2.0641 | 2016-09-20 09:08 | 2.0677 | 2.0849 | 2016-09-20 09:10 | 2.0542 | 2.071
65.7 | 37.6 | 125.1 | 20 | 9 | 12 | 0 | 9 | 16 | C | 83 | IZO | 61.0229 | 34.5048 | 111.7482 | 264.2663 | 255.5428 | 264.3753 | 271.0472 | 266.5955 | 0.49473 | 0.21842 | 0.49788 | 0.092151 | 0.50102 | 0.47259 | 2016-09-20 09:12 | 2.0213 | 2.0371 | 2016-09-20 09:13 | 2.0085 | 2.024 | 2016-09-20 09:14 | 1.9959 | 2.011 | 2016-09-20 09:12 | 2.0149 | 2.0305 | 2016-09-20 09:14 | 2.0022 | 2.0174
63.9 | 36.6 | 122.1 | 20 | 9 | 16 | 0 | 9 | 16 | C | 83 | IZO | 59.3592 | 33.5113 | 108.9092 | 263.3717 | 253.5704 | 264.0444 | 270.9847 | 266.6801 | 0.50727 | 0.21768 | 0.51039 | 0.09144 | 0.51349 | 0.472 | 2016-09-20 09:16 | 1.9713 | 1.9857 | 2016-09-20 09:17 | 1.9593 | 1.9733 | 2016-09-20 09:18 | 1.9475 | 1.9611 | 2016-09-20 09:16 | 1.9653 | 1.9794 | 2016-09-20 09:18 | 1.9534 | 1.9672
62.6 | 35.9 | 119.3 | 20 | 9 | 20 | 0 | 9 | 16 | C | 83 | IZO | 58.1562 | 32.8141 | 106.2794 | 264.7679 | 255.0661 | 263.8466 | 272.3021 | 266.055 | 0.51968 | 0.21883 | 0.52276 | 0.09198 | 0.52583 | 0.47165 | 2016-09-20 09:20 | 1.9243 | 1.9373 | 2016-09-20 09:21 | 1.9129 | 1.9256 | 2016-09-20 09:22 | 1.9018 | 1.9142 | 2016-09-20 09:20 | 1.9186 | 1.9314 | 2016-09-20 09:22 | 1.9073 | 1.9199
61.2 | 35.1 | 116.7 | 20 | 9 | 24 | 0 | 9 | 16 | C | 83 | IZO | 56.8588 | 32.0155 | 103.84 | 265.0924 | 254.443 | 263.7725 | 273.3585 | 266.1202 | 0.53194 | 0.2191 | 0.53498 | 0.091755 | 0.53801 | 0.47151 | 2016-09-20 09:24 | 1.8799 | 1.8918 | 2016-09-20 09:25 | 1.8692 | 1.8808 | 2016-09-20 09:26 | 1.8587 | 1.87 | 2016-09-20 09:24 | 1.8746 | 1.8863 | 2016-09-20 09:26 | 1.8639 | 1.8754
59.6 | 34.4 | 114.2 | 20 | 9 | 28 | 0 | 9 | 16 | C | 83 | IZO | 55.373 | 31.3155 | 101.4897 | 263.6877 | 254.6355 | 263.5529 | 270.7159 | 265.7969 | 0.54405 | 0.21794 | 0.54706 | 0.091824 | 0.55005 | 0.47112 | 2016-09-20 09:28 | 1.8381 | 1.8489 | 2016-09-20 09:29 | 1.828 | 1.8385 | 2016-09-20 09:30 | 1.818 | 1.8283 | 2016-09-20 09:28 | 1.833 | 1.8437 | 2016-09-20 09:30 | 1.823 | 1.8334
58.3 | 33.5 | 111.9 | 20 | 9 | 32 | 0 | 9 | 16 | C | 83 | IZO | 54.1628 | 30.4138 | 99.3294 | 263.5995 | 251.0962 | 263.5265 | 273.2976 | 266.6587 | 0.55601 | 0.21787 | 0.55898 | 0.090548 | 0.56193 | 0.47107 | 2016-09-20 09:32 | 1.7985 | 1.8084 | 2016-09-20 09:33 | 1.789 | 1.7986 | 2016-09-20 09:34 | 1.7796 | 1.789 | 2016-09-20 09:32 | 1.7937 | 1.8035 | 2016-09-20 09:34 | 1.7843 | 1.7938
57.1 | 32.9 | 109.4 | 20 | 9 | 36 | 0 | 9 | 16 | C | 83 | IZO | 53.0446 | 29.8117 | 96.9912 | 263.6888 | 251.5666 | 262.5543 | 273.0907 | 265.3224 | 0.56782 | 0.21794 | 0.57074 | 0.090718 | 0.57366 | 0.46934 | 2016-09-20 09:36 | 1.7611 | 1.7701 | 2016-09-20 09:37 | 1.7521 | 1.7609 | 2016-09-20 09:38 | 1.7432 | 1.7518 | 2016-09-20 09:36 | 1.7566 | 1.7655 | 2016-09-20 09:38 | 1.7476 | 1.7563
56.1 | 32.4 | 107.5 | 20 | 9 | 40 | 0 | 9 | 16 | C | 83 | IZO | 52.1114 | 29.3093 | 95.2233 | 264.6523 | 253.117 | 263.0547 | 273.5989 | 265.5578 | 0.57946 | 0.21874 | 0.58234 | 0.091277 | 0.58522 | 0.47023 | 2016-09-20 09:40 | 1.7258 | 1.734 | 2016-09-20 09:41 | 1.7172 | 1.7253 | 2016-09-20 09:42 | 1.7088 | 1.7166 | 2016-09-20 09:40 | 1.7215 | 1.7296 | 2016-09-20 09:42 | 1.713 | 1.7209
54.9 | 31.8 | 105.6 | 20 | 9 | 44 | 0 | 9 | 16 | C | 83 | IZO | 50.9893 | 28.7056 | 93.4563 | 263.897 | 252.5376 | 263.2407 | 272.7064 | 265.9375 | 0.59093 | 0.21811 | 0.59378 | 0.091068 | 0.59661 | 0.47056 | 2016-09-20 09:44 | 1.6922 | 1.6998 | 2016-09-20 09:45 | 1.6841 | 1.6915 | 2016-09-20 09:46 | 1.6761 | 1.6834 | 2016-09-20 09:44 | 1.6882 | 1.6956 | 2016-09-20 09:46 | 1.6801 | 1.6874
53.9 | 31.3 | 103.8 | 20 | 9 | 48 | 0 | 9 | 16 | C | 83 | IZO | 50.0517 | 28.2017 | 91.7764 | 264.0698 | 253.1136 | 263.4086 | 272.5663 | 266.0025 | 0.60224 | 0.21825 | 0.60504 | 0.091276 | 0.60783 | 0.47086 | 2016-09-20 09:48 | 1.6605 | 1.6674 | 2016-09-20 09:49 | 1.6528 | 1.6595 | 2016-09-20 09:50 | 1.6452 | 1.6518 | 2016-09-20 09:48 | 1.6566 | 1.6635 | 2016-09-20 09:50 | 1.649 | 1.6557
47.3 | 27.7 | 92.1 | 20 | 10 | 20 | 0 | 9 | 16 | C | 83 | IZO | 43.7978 | 24.5504 | 80.7746 | 263.1379 | 249.0911 | 263.6542 | 274.0215 | 267.3283 | 0.68615 | 0.21748 | 0.68857 | 0.089825 | 0.69098 | 0.4713 | 2016-09-20 10:20 | 1.4574 | 1.461 | 2016-09-20 10:21 | 1.4523 | 1.4558 | 2016-09-20 10:22 | 1.4472 | 1.4506 | 2016-09-20 10:20 | 1.4548 | 1.4583 | 2016-09-20 10:22 | 1.4497 | 1.4532
46.7 | 27.7 | 90.9 | 20 | 10 | 24 | 0 | 9 | 16 | C | 83 | IZO | 43.2223 | 24.5504 | 79.6511 | 263.4048 | 255.6378 | 263.5584 | 269.4272 | 265.5541 | 0.69577 | 0.2177 | 0.69814 | 0.092186 | 0.7005 | 0.47113 | 2016-09-20 10:24 | 1.4373 | 1.4405 | 2016-09-20 10:25 | 1.4324 | 1.4356 | 2016-09-20 10:26 | 1.4276 | 1.4307 | 2016-09-20 10:24 | 1.4348 | 1.4381 | 2016-09-20 10:25 | 1.43 | 1.4331
46.1 | 27.3 | 89.8 | 20 | 10 | 28 | 0 | 9 | 16 | C | 83 | IZO | 42.645 | 24.1418 | 78.6247 | 263.417 | 254.0263 | 263.6374 | 270.6958 | 266.0605 | 0.70518 | 0.21771 | 0.7075 | 0.091605 | 0.70981 | 0.47127 | 2016-09-20 10:28 | 1.4181 | 1.4211 | 2016-09-20 10:29 | 1.4134 | 1.4164 | 2016-09-20 10:30 | 1.4088 | 1.4117 | 2016-09-20 10:28 | 1.4157 | 1.4187 | 2016-09-20 10:30 | 1.4111 | 1.414
45.5 | 27.1 | 88.7 | 20 | 10 | 32 | 0 | 9 | 16 | C | 83 | IZO | 42.0665 | 23.9372 | 77.6008 | 263.1796 | 256.119 | 263.5284 | 268.6544 | 265.3951 | 0.71438 | 0.21752 | 0.71665 | 0.092359 | 0.71891 | 0.47108 | 2016-09-20 10:32 | 1.3998 | 1.4026 | 2016-09-20 10:33 | 1.3954 | 1.3981 | 2016-09-20 10:34 | 1.391 | 1.3937 | 2016-09-20 10:32 | 1.3976 | 1.4003 | 2016-09-20 10:34 | 1.3932 | 1.3959
44.9 | 26.7 | 87.8 | 20 | 10 | 36 | 0 | 9 | 16 | C | 83 | IZO | 41.4868 | 23.5278 | 76.7645 | 262.6964 | 253.8435 | 263.9878 | 269.5582 | 266.5458 | 0.72338 | 0.21712 | 0.72559 | 0.091539 | 0.72779 | 0.4719 | 2016-09-20 10:36 | 1.3824 | 1.385 | 2016-09-20 10:37 | 1.3782 | 1.3807 | 2016-09-20 10:38 | 1.374 | 1.3765 | 2016-09-20 10:36 | 1.3803 | 1.3828 | 2016-09-20 10:38 | 1.3761 | 1.3786
44.6 | 26.5 | 86.8 | 20 | 10 | 40 | 0 | 9 | 16 | C | 83 | IZO | 41.1965 | 23.3229 | 75.8366 | 264.544 | 255.3903 | 263.8947 | 271.6383 | 266.0384 | 0.73215 | 0.21865 | 0.73432 | 0.092096 | 0.73646 | 0.47173 | 2016-09-20 10:40 | 1.3658 | 1.3682 | 2016-09-20 10:41 | 1.3618 | 1.3641 | 2016-09-20 10:42 | 1.3578 | 1.3601 | 2016-09-20 10:40 | 1.3638 | 1.3662 | 2016-09-20 10:41 | 1.3598 | 1.3621
44.1 | 26.6 | 86.1 | 20 | 10 | 44 | 0 | 9 | 16 | C | 83 | IZO | 40.7119 | 23.4253 | 75.1877 | 264.4809 | 263.0296 | 264.7823 | 265.613 | 265.2199 | 0.74072 | 0.21859 | 0.74282 | 0.094851 | 0.74492 | 0.47332 | 2016-09-20 10:44 | 1.35 | 1.3522 | 2016-09-20 10:45 | 1.3462 | 1.3483 | 2016-09-20 10:46 | 1.3424 | 1.3445 | 2016-09-20 10:44 | 1.3481 | 1.3503 | 2016-09-20 10:46 | 1.3443 | 1.3464
43.7 | 26.2 | 85.1 | 20 | 10 | 48 | 0 | 9 | 16 | C | 83 | IZO | 40.3235 | 23.0152 | 74.2611 | 265.0822 | 259.8948 | 264.3467 | 269.106 | 265.4666 | 0.74906 | 0.21909 | 0.75111 | 0.093721 | 0.75315 | 0.47254 | 2016-09-20 10:48 | 1.335 | 1.337 | 2016-09-20 10:49 | 1.3314 | 1.3333 | 2016-09-20 10:50 | 1.3278 | 1.3297 | 2016-09-20 10:48 | 1.3332 | 1.3352 | 2016-09-20 10:50 | 1.3296 | 1.3315
43.2 | 26 | 84.3 | 20 | 10 | 52 | 0 | 9 | 16 | C | 83 | IZO | 39.8372 | 22.8099 | 73.5204 | 264.5999 | 260.7435 | 264.5251 | 267.593 | 265.4757 | 0.75718 | 0.21869 | 0.75918 | 0.094027 | 0.76116 | 0.47286 | 2016-09-20 10:52 | 1.3207 | 1.3225 | 2016-09-20 10:53 | 1.3172 | 1.319 | 2016-09-20 10:54 | 1.3138 | 1.3156 | 2016-09-20 10:52 | 1.3189 | 1.3208 | 2016-09-20 10:54 | 1.3155 | 1.3173
42.7 | 25.7 | 83.4 | 20 | 10 | 56 | 0 | 9 | 16 | C | 83 | IZO | 39.3499 | 22.5018 | 72.6888 | 263.905 | 259.1693 | 264.1538 | 267.5785 | 265.4083 | 0.76508 | 0.21812 | 0.76702 | 0.093459 | 0.76894 | 0.4722 | 2016-09-20 10:56 | 1.3071 | 1.3088 | 2016-09-20 10:57 | 1.3037 | 1.3054 | 2016-09-20 10:58 | 1.3005 | 1.3021 | 2016-09-20 10:56 | 1.3054 | 1.3071 | 2016-09-20 10:57 | 1.3021 | 1.3038
42.3 | 25.4 | 82.5 | 20 | 11 | 0 | 0 | 9 | 16 | C | 83 | IZO | 38.9593 | 22.1933 | 71.8585 | 263.9139 | 257.3133 | 263.6236 | 269.0307 | 265.2133 | 0.77275 | 0.21813 | 0.77463 | 0.09279 | 0.7765 | 0.47125 | 2016-09-20 11:00 | 1.2941 | 1.2956 | 2016-09-20 11:01 | 1.2909 | 1.2925 | 2016-09-20 11:02 | 1.2878 | 1.2893 | 2016-09-20 11:01 | 1.2925 | 1.2941 | 2016-09-20 11:02 | 1.2894 | 1.2909
41.9 | 25.1 | 81.9 | 20 | 11 | 4 | 0 | 9 | 16 | C | 83 | IZO | 38.568 | 21.8845 | 71.3054 | 263.7362 | 255.1798 | 264.1469 | 270.3666 | 266.4081 | 0.78019 | 0.21798 | 0.78202 | 0.092021 | 0.78383 | 0.47218 | 2016-09-20 11:04 | 1.2817 | 1.2832 | 2016-09-20 11:05 | 1.2787 | 1.2802 | 2016-09-20 11:06 | 1.2758 | 1.2772 | 2016-09-20 11:04 | 1.2802 | 1.2817 | 2016-09-20 11:06 | 1.2773 | 1.2787
41.6 | 25 | 81.2 | 20 | 11 | 8 | 0 | 9 | 16 | C | 83 | IZO | 38.2741 | 21.7816 | 70.6602 | 264.3079 | 257.2829 | 264.1252 | 269.7528 | 265.8495 | 0.7874 | 0.21845 | 0.78917 | 0.092779 | 0.79092 | 0.47214 | 2016-09-20 11:08 | 1.27 | 1.2713 | 2016-09-20 11:09 | 1.2672 | 1.2685 | 2016-09-20 11:10 | 1.2644 | 1.2656 | 2016-09-20 11:08 | 1.2686 | 1.2699 | 2016-09-20 11:10 | 1.2658 | 1.267
41.5 | 24.8 | 80.6 | 20 | 11 | 12 | 0 | 9 | 16 | C | 83 | IZO | 38.176 | 21.5755 | 70.1072 | 266.6027 | 256.9273 | 264.3708 | 274.0989 | 266.2471 | 0.79437 | 0.22035 | 0.79608 | 0.092651 | 0.79777 | 0.47258 | 2016-09-20 11:12 | 1.2589 | 1.2601 | 2016-09-20 11:13 | 1.2562 | 1.2573 | 2016-09-20 11:14 | 1.2535 | 1.2547 | 2016-09-20 11:12 | 1.2575 | 1.2587 | 2016-09-20 11:13 | 1.2548 | 1.256
41 | 24.8 | 80 | 20 | 11 | 16 | 0 | 9 | 16 | C | 83 | IZO | 37.685 | 21.5755 | 69.5539 | 264.9628 | 260.9384 | 264.4802 | 268.0846 | 265.3706 | 0.80111 | 0.21899 | 0.80276 | 0.094097 | 0.80439 | 0.47278 | 2016-09-20 11:16 | 1.2483 | 1.2494 | 2016-09-20 11:17 | 1.2457 | 1.2468 | 2016-09-20 11:18 | 1.2432 | 1.2442 | 2016-09-20 11:17 | 1.247 | 1.2481 | 2016-09-20 11:18 | 1.2444 | 1.2455
40.7 | 24.6 | 79.4 | 20 | 11 | 20 | 0 | 9 | 16 | C | 83 | IZO | 37.3898 | 21.3693 | 69.0002 | 265.0486 | 260.1787 | 264.4539 | 268.8248 | 265.5297 | 0.80761 | 0.21906 | 0.8092 | 0.093823 | 0.81077 | 0.47273 | 2016-09-20 11:20 | 1.2382 | 1.2392 | 2016-09-20 11:21 | 1.2358 | 1.2368 | 2016-09-20 11:22 | 1.2334 | 1.2344 | 2016-09-20 11:20 | 1.237 | 1.238 | 2016-09-20 11:22 | 1.2346 | 1.2356
40.4 | 24.4 | 78.9 | 20 | 11 | 24 | 0 | 9 | 16 | C | 83 | IZO | 37.0942 | 21.163 | 68.5382 | 264.9753 | 259.2 | 264.715 | 269.4521 | 266.104 | 0.81387 | 0.219 | 0.8154 | 0.09347 | 0.81691 | 0.4732 | 2016-09-20 11:24 | 1.2287 | 1.2296 | 2016-09-20 11:25 | 1.2264 | 1.2273 | 2016-09-20 11:26 | 1.2241 | 1.225 | 2016-09-20 11:24 | 1.2275 | 1.2285 | 2016-09-20 11:26 | 1.2253 | 1.2262

## param
-0.9 | 1e-05 |  | -1.9 | 0
0 | 1 |  | -1 | 1
0.9 | 1e-05 |  | 1 | 1
|  |  | 1.9 | 0
304.6 | 1e-07 | 310.55 | 1e-06 | 316.7 | 1e-06 | 323.5 | 1e-06 | 323.1 | 1e-06 | 330.5 | 1e-06 | 338 | 1e-06
305.5 | 1000 | 311.45 | 1000 | 317.6 | 1000 | 324.4 | 1000 | 324 | 1000 | 331.4 | 1000 | 338.9 | 1000
306.4 | 1e-07 | 312.35 | 1e-06 | 318.5 | 1e-06 | 326.4 | 1000 | 326 | 1000 | 333.4 | 1000 | 340.9 | 1000
[6 empty] | 327.3 | 1e-06 | 326.9 | 1e-06 | 334.3 | 1e-06 | 341.8 | 1e-06
1.1111
304.6 | 0.001 | 310.55 | 0.001 | 316.7 | 0.001 | 323.5 | 0 | 323.1 | 0.001 | 330.5 | 0.001 | 338 | 1e-07
305.5 | 100 | 311.45 | 100 | 317.6 | 100 | 324.4 | 100 | 324 | 100 | 331.4 | 100 | 338.9 | 100
306.4 | 0.001 | 312.35 | 0.001 | 318.5 | 0.001 | 326.4 | 100 | 326 | 100 | 333.4 | 100 | 340.9 | 100
[6 empty] | 327.3 | 0 | 326.9 | 0.001 | 334.3 | 0.001 | 341.8 | 1e-07
58 | 0
92 | 0
code | number | name | Dobson | earth_radius | LONG | Hemisphere | degs | min | sec | LAT | Hemisphere | deg | min | sec | TimeZone | stn Pressure | elevation | Ozone_Height | LongRadian | LatRadian
IZO | 999 | Izana Observatory,Spain | 83 | 6373593 | 16.4973 | W | 16 | 29 | 25 | 28.2986 | N | 28 | 17 | 55 | 0 | 770 | 2360 | 26000 | 0.28781 | 0.4939
=IF( C13<90,(earth_radius+Ozone_height)/(SQRT(((earth_radius+Ozone_height)^2)-((earth_radius+Elevation)^2)*(SIN(RADIANS(C13)))^2)),-90000)
Public Function AirMass#(CosZenAng As Double) | [14 empty] | p0 | 0.75993
AirMass# = 1# / (CosZenAng + 0.15 / (90 - RadsToDegs#(ArcCos#(CosZenAng)) + 3.885) ^ 1.253)
End Function
Ozone Absorption and Molecular Scattering coefficients for Use
with Dobson ozone Spectrophotometers Beginning 1 January 1992
Wavelength | ALFA |  | BETA | d(BETA)/d(ALFA) | [5 empty] | IUP Corr-factor | B&BP Corr_factor | D&M |  |  | α_Komhyr r adj
(nm) | (atm-cm)-1 |  | (atm)-1 | (atm-cm/atm) |  |  |  |  | A_ pair | 0.9898 | 0.9969 | 0.9889 |  |  | 1.806
[9 empty] | C_pair | 0.9922 | 0.9983 | 0.9845 |  |  | 0.833
305 | 0.489 | [7 empty] | D_pair | 0.9642 | 0.9825 | 0.9562 |  |  | 0.374
325 | 0.375 | [7 empty] | AD_pair | 0.9965 | 1.0006 | 0.9975 |  |  | 1.432
Alfa_A | 1.7876 | Beta_A | 114 | 63 | 1000 |  |  |  | CD_pair | 1.015 | 1.0112 | 1.0076 |  |  | 0.459
|  |  | 0 | 0
308.9 | 0.466 |  | 0 | 0
329.1 | 0.355 |  | 0 | 0
Alfa_B | 1.192 | Beta_B | 111 | 93
|  |  | 0 | 0
311.5 | 0.45 |  | 0 | 0 |  |  | 5
332.4 | 0.341 |  | 0 | 0
Alfa_C | 0.8265 | Beta_C | 109 | 131
|  |  | 0 | 0
317.5 | 0.414 |  | 0 | 0
339.9 | 0.31 |  | 0 | 0
Alfa_D | 0.36061 | Beta_D | 104 | 278
|  |  | 0 | 0
Alfa_AD | 1.427 | Beta_AD | 10 | 7
Alfa_BD | 0.818 | Beta_BD | 7 | 9
Alfa_CD | 0.46588 | Beta_CD | 5 | 11

## Nfunc083
D83RA | D83NA | D83RC | D83NC | D83RD | D83ND | Rby5 | a | c | d | in pretty print
5 | -1.51 | 5 | 2.18 | 5 | 3.26 | 5 | -1.51 | 2.18 | 3.26 | -1.5 | 2.2 | 3.3 | 0.01 | 0.02 | 0.04
10.74 | 4.7 | 10.76 | 8.39 | 10.72 | 9.62 | 15 | 9.2606 | 12.8971 | 14.1824 | 9.2 | 12.9 | 14.2 | -0.060647 | 0.0029402 | 0.017648
16.48 | 10.83 | 16.51 | 14.49 | 16.44 | 15.7 | 25 | 19.7018 | 23.3471 | 24.6275 | 19.7 | 23.3 | 24.6 | -0.001759 | -0.047139 | -0.027526
22.22 | 16.83 | 22.27 | 20.51 | 22.16 | 21.69 | 35 | 29.9065 | 33.5082 | 34.7616 | 29.9 | 33.5 | 34.8 | -0.0065067 | -0.0081693 | 0.038446
27.96 | 22.74 | 28.02 | 26.46 | 27.88 | 27.58 | 45 | 39.8449 | 43.394 | 44.5867 | 39.8 | 43.4 | 44.6 | -0.044871 | 0.0060056 | 0.013291
33.7 | 28.59 | 33.78 | 32.28 | 33.6 | 33.36 | 55 | 49.3404 | 52.8934 | 54.076 | 49.3 | 52.9 | 54.1 | -0.040402 | 0.0065573 | 0.02402
39.87 | 34.8 | 39.99 | 38.5 | 40.01 | 39.73 | 65 | 58.6907 | 62.1867 | 63.3142 | 58.7 | 62.2 | 63.3 | 0.0093361 | 0.013335 | -0.014152
46.12 | 40.93 | 46.25 | 44.6 | 46.26 | 45.8 | 75 | 67.9965 | 71.4015 | 72.4459 | 68 | 71.4 | 72.4 | 0.0034853 | -0.0015358 | -0.045888
52.43 | 46.93 | 52.58 | 50.62 | 52.56 | 51.79 | 85 | 77.243 | 80.6092 | 81.6241 | 77.2 | 80.6 | 81.6 | -0.043043 | -0.0091746 | -0.024141
58.76 | 52.85 | 58.93 | 56.56 | 58.88 | 57.68 | 95 | 86.5778 | 89.8754 | 90.8361 | 86.6 | 89.9 | 90.8 | 0.022224 | 0.024624 | -0.036083
65.01 | 58.7 | 65.22 | 62.39 | 65.17 | 63.47 | 105 | 95.9718 | 99.1641 | 100.0591 | 96 | 99.2 | 100.1 | 0.028156 | 0.03591 | 0.040888
71.68 | 64.91 | 71.96 | 68.6 | 72.14 | 69.83 | 115 | 105.3164 | 108.4291 | 109.276 | 105.3 | 108.4 | 109.3 | -0.016442 | -0.029128 | 0.024023
78.27 | 71.03 | 78.58 | 74.7 | 78.77 | 75.9 | 125 | 114.7278 | 117.7197 | 118.4987 | 114.7 | 117.7 | 118.5 | -0.027761 | -0.019725 | 0.0013422
84.77 | 77.03 | 85.12 | 80.72 | 85.3 | 81.9 | 135 | 124.2697 | 127.0863 | 127.7279 | 124.3 | 127.1 | 127.7 | 0.030284 | 0.013723 | -0.027928
91.14 | 82.95 | 91.54 | 86.67 | 91.69 | 87.78 | 145 | 133.8086 | 136.5037 | 137.0681 | 133.8 | 136.5 | 137.1 | -0.0086133 | -0.0037106 | 0.031879
97.36 | 88.8 | 97.82 | 92.49 | 97.96 | 93.57 | 155 | 143.4602 | 145.9931 | 146.5049 | 143.5 | 146 | 146.5 | 0.03978 | 0.0069477 | -0.0049043
103.97 | 95.01 | 104.5 | 98.7 | 104.86 | 99.93 | 165 | 153.1836 | 155.5627 | 155.9752 | 153.2 | 155.6 | 156 | 0.016436 | 0.037317 | 0.02477
110.54 | 101.13 | 111.08 | 104.8 | 111.45 | 106.01 | 175 | 162.9608 | 165.2327 | 165.5125 | 163 | 165.2 | 165.5 | 0.039212 | -0.032674 | -0.012494
116.94 | 107.14 | 117.58 | 110.82 | 117.96 | 112 | 185 | 172.7866 | 174.9334 | 175.074 | 172.8 | 174.9 | 175.1 | 0.013394 | -0.033433 | 0.026028
123.23 | 113.05 | 123.98 | 116.77 | 124.34 | 117.89 | 195 | 182.6394 | 184.6057 | 184.6243 | 182.6 | 184.6 | 184.6 | -0.039352 | -0.0057191 | -0.024324
129.38 | 118.9 | 130.22 | 122.59 | 130.61 | 123.67 | 205 | 192.4841 | 194.2836 | 194.1583 | 192.5 | 194.3 | 194.2 | 0.015876 | 0.016413 | 0.041746
135.88 | 125.11 | 136.83 | 128.81 | 137.48 | 130.03 | 215 | 202.4109 | 204.0518 | 203.7711 | 202.4 | 204.1 | 203.8 | -0.010885 | 0.048175 | 0.028851
142.31 | 131.24 | 143.3 | 134.9 | 143.98 | 136.11 | 225 | 212.4107 | 213.9099 | 213.4697 | 212.4 | 213.9 | 213.5 | -0.010729 | -0.0099222 | 0.030315
148.58 | 137.24 | 149.67 | 140.92 | 150.34 | 142.1 | 235 | 222.6125 | 223.9066 | 223.3561 | 222.6 | 223.9 | 223.4 | -0.01248 | -0.0066373 | 0.043946
154.69 | 143.16 | 155.92 | 146.87 | 156.57 | 147.99 | 245 | 232.9094 | 234.043 | 233.3305 | 232.9 | 234 | 233.3 | -0.0094379 | -0.04302 | -0.030472
160.72 | 149 | 162.02 | 152.69 | 162.68 | 153.77 | 255 | 243.0375 | 244.1782 | 243.3484 | 243 | 244.2 | 243.3 | -0.037453 | 0.021839 | -0.048358
167.07 | 155.21 | 168.46 | 158.91 | 169.37 | 160.14 | 265 | 253.2588 | 254.2376 | 253.357 | 253.3 | 254.2 | 253.4 | 0.041246 | -0.037564 | 0.043032
173.34 | 161.34 | 174.77 | 165.01 | 175.73 | 166.21 | 275 | 263.5293 | 264.3569 | 263.3423 | 263.5 | 264.4 | 263.3 | -0.029307 | 0.043099 | -0.042337
179.48 | 167.34 | 180.98 | 171.03 | 181.99 | 172.2 | 285 | 273.6505 | 274.5418 | 273.4295 | 273.7 | 274.5 | 273.4 | 0.049496 | -0.041795 | -0.029511
185.48 | 173.26 | 187.11 | 176.98 | 188.16 | 178.09 |  |  |  |  | 284 | 284.8 | 283.6 | -0.080761 | -0.032459 | 0.10211
191.42 | 179.11 | 193.13 | 182.8 | 194.22 | 183.88
197.72 | 185.32 | 199.56 | 189.01 | 200.89 | 190.24
203.94 | 191.44 | 205.85 | 195.11 | 207.25 | 196.31
210.02 | 197.44 | 212.02 | 201.13 | 213.48 | 202.31
215.95 | 203.36 | 218.08 | 207.08 | 219.6 | 208.2
221.81 | 209.21 | 223.98 | 212.9 | 225.52 | 213.98
227.98 | 215.42 | 230.23 | 219.11 | 231.96 | 220.34
233.97 | 221.55 | 236.29 | 225.21 | 238.08 | 226.42
239.77 | 227.55 | 242.22 | 231.23 | 244.08 | 232.41
245.54 | 233.46 | 248.1 | 237.18 | 249.96 | 238.3
251.31 | 239.31 | 253.83 | 243 | 255.73 | 244.08
257.45 | 245.52 | 260.03 | 249.22 | 262.08 | 250.45
263.44 | 251.65 | 266.07 | 255.32 | 268.18 | 256.52
269.25 | 257.65 | 272.01 | 261.34 | 274.17 | 262.51
275.04 | 263.57 | 277.89 | 267.28 | 280.03 | 268.4
280.82 | 269.42 | 283.59 | 273.11 | 285.75 | 274.19
286.94 | 275.63 | 289.71 | 279.32 | 292.01 | 280.55
292.84 | 281.75 | 295.61 | 285.42 | 298.01 | 286.62
298.49 | 287.75

## comp page
Komyhr adjusted |  |  |  |  | IUP multiply |  |  |  |  | B&P B multiply |  |  |  |  | D&M multiply
268.3915 | 275.7695 | 267.9614 | 262.4369 | 265.9041 | 270.5014 | 286.0086 | 270.7227 | 258.5586 | 266.838 | 268.8486 | 280.6814 | 268.7946 | 259.5302 | 265.7446 | 272.6171 | 288.4015 | 270.9691 | 260.4575 | 266.5705
D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD
264.7325 | 266.4717 | 264.3446 | 263.652 | 263.6841 | 266.8137 | 276.3656 | 267.0687 | 259.7557 | 264.6102 | 265.1833 | 271.218 | 265.1666 | 260.7318 | 263.526 | 268.9005 | 278.6778 | 267.3117 | 261.6634 | 264.345
268.5041 | 267.2848 | 265.1969 | 269.7742 | 264.565 | 270.6149 | 277.2089 | 267.9297 | 265.7873 | 265.4943 | 268.9613 | 272.0456 | 266.0215 | 266.7861 | 264.4064 | 272.7314 | 279.5282 | 268.1736 | 267.7393 | 265.2281
265.1548 | 267.7076 | 265.8913 | 263.3046 | 265.3448 | 267.2392 | 277.6474 | 268.6313 | 259.4134 | 266.2768 | 265.6063 | 272.4759 | 266.7181 | 260.3883 | 265.1857 | 269.3294 | 279.9703 | 268.8758 | 261.3186 | 266.0099
264.8362 | 268.2334 | 266.343 | 262.2611 | 265.7885 | 266.9181 | 278.1927 | 269.0877 | 258.3854 | 266.7221 | 265.2871 | 273.0111 | 267.1712 | 259.3563 | 265.6292 | 269.0057 | 280.5202 | 269.3326 | 260.283 | 266.4547
265.1375 | 269.3693 | 266.1339 | 261.8531 | 265.2373 | 267.2219 | 279.3707 | 268.8764 | 257.9834 | 266.1689 | 265.589 | 274.1672 | 266.9614 | 258.9529 | 265.0783 | 269.3119 | 281.7081 | 269.1211 | 259.8781 | 265.9021
266.1345 | 268.5326 | 266.1113 | 264.3204 | 265.4348 | 268.2267 | 278.503 | 268.8536 | 260.4142 | 266.3671 | 266.5877 | 273.3157 | 266.9388 | 261.3928 | 265.2756 | 270.3245 | 280.8331 | 269.0983 | 262.3267 | 266.1
266.7449 | 270.4133 | 266.5675 | 263.8762 | 265.525 | 268.8418 | 280.4535 | 269.3145 | 259.9766 | 266.4576 | 267.1991 | 275.2298 | 267.3965 | 260.9536 | 265.3658 | 270.9445 | 282.7999 | 269.5596 | 261.8859 | 266.1905
266.0513 | 271.7368 | 267.1369 | 261.5232 | 265.9024 | 268.1428 | 281.8261 | 269.8897 | 257.6583 | 266.8363 | 266.5044 | 276.5769 | 267.9676 | 258.6265 | 265.743 | 270.24 | 284.184 | 270.1354 | 259.5506 | 266.5688
266.7401 | 271.0499 | 266.683 | 263.3196 | 265.5135 | 268.8371 | 281.1138 | 269.4312 | 259.4282 | 266.4461 | 267.1944 | 275.8777 | 267.5122 | 260.4031 | 265.3543 | 270.9397 | 283.4657 | 269.6764 | 261.3335 | 266.179
266.9044 | 269.1747 | 266.7246 | 265.1345 | 266.0593 | 269.0026 | 279.169 | 269.4732 | 261.2162 | 266.9938 | 267.3589 | 273.9692 | 267.554 | 262.1978 | 265.8998 | 271.1065 | 281.5046 | 269.7185 | 263.1346 | 266.7261
266.5945 | 270.802 | 266.5124 | 263.2368 | 265.3696 | 268.6903 | 280.8566 | 269.2588 | 259.3466 | 266.3017 | 267.0485 | 275.6254 | 267.3412 | 260.3212 | 265.2105 | 270.7918 | 283.2064 | 269.5039 | 261.2513 | 266.0347
266.7337 | 269.91 | 267.3771 | 264.2082 | 266.6957 | 268.8306 | 279.9315 | 270.1325 | 260.3037 | 267.6324 | 267.188 | 274.7176 | 268.2086 | 261.2819 | 266.5357 | 270.9332 | 282.2736 | 270.3783 | 262.2154 | 267.3641
266.7017 | 270.3913 | 267.0774 | 263.7512 | 266.1941 | 268.7984 | 280.4307 | 269.8297 | 259.8534 | 267.1291 | 267.1559 | 275.2074 | 267.9079 | 260.8299 | 266.0345 | 270.9007 | 282.7769 | 270.0752 | 261.7618 | 266.8613
266.5923 | 271.5517 | 267.021 | 262.6011 | 265.8218 | 268.688 | 281.6342 | 269.7727 | 258.7203 | 266.7555 | 267.0462 | 276.3885 | 267.8514 | 259.6926 | 265.6624 | 270.7895 | 283.9904 | 270.0182 | 260.6204 | 266.488
267.3632 | 271.461 | 266.6941 | 264.0689 | 265.4348 | 269.465 | 281.5402 | 269.4424 | 260.1664 | 266.3671 | 267.8184 | 276.2962 | 267.5234 | 261.1441 | 265.2756 | 271.5725 | 283.8957 | 269.6876 | 262.0771 | 266.1001
267.2834 | 272.369 | 267.3354 | 263.1799 | 266.0079 | 269.3846 | 282.4818 | 270.0904 | 259.2905 | 266.9422 | 267.7385 | 277.2203 | 268.1668 | 260.2649 | 265.8484 | 271.4915 | 284.8452 | 270.3362 | 261.1948 | 266.6746
267.8399 | 272.1793 | 267.3553 | 264.3404 | 266.0838 | 269.9454 | 282.2851 | 270.1105 | 260.4339 | 267.0184 | 268.296 | 277.0273 | 268.1867 | 261.4126 | 265.9242 | 272.0567 | 284.6469 | 270.3563 | 262.3466 | 266.7507
267.79 | 272.2267 | 267.4763 | 264.205 | 266.2261 | 269.8952 | 282.3342 | 270.2327 | 260.3005 | 267.1611 | 268.246 | 277.0755 | 268.3081 | 261.2787 | 266.0664 | 272.0061 | 284.6964 | 270.4787 | 262.2122 | 266.8933
267.2456 | 271.4126 | 267.7117 | 263.8777 | 266.7363 | 269.3465 | 281.49 | 270.4705 | 259.978 | 267.6732 | 267.7007 | 276.247 | 268.5442 | 260.955 | 266.5764 | 271.4531 | 283.845 | 270.7167 | 261.8873 | 267.4048
268.2065 | 272.2995 | 267.7112 | 264.8966 | 266.5049 | 270.3149 | 282.4097 | 270.47 | 260.9818 | 267.4409 | 268.6632 | 277.1496 | 268.5437 | 261.9626 | 266.345 | 272.4291 | 284.7725 | 270.7162 | 262.8985 | 267.1728
267.471 | 271.0485 | 267.5339 | 264.579 | 266.6086 | 269.5737 | 281.1124 | 270.2909 | 260.6689 | 267.545 | 267.9265 | 275.8764 | 268.3658 | 261.6485 | 266.4487 | 271.6821 | 283.4643 | 270.5368 | 262.5833 | 267.2768
269.0068 | 273.1343 | 267.4866 | 265.6648 | 266.0048 | 271.1216 | 283.2756 | 270.2431 | 261.7387 | 266.9391 | 269.4649 | 277.9993 | 268.3184 | 262.7223 | 265.8453 | 273.2421 | 285.6456 | 270.4891 | 263.661 | 266.6715
269.4775 | 273.1696 | 267.5601 | 266.4886 | 266.0888 | 271.596 | 283.3121 | 270.3173 | 262.5504 | 267.0234 | 269.9364 | 278.0352 | 268.3921 | 263.537 | 265.9293 | 273.7202 | 285.6825 | 270.5634 | 264.4786 | 266.7557
268.2307 | 273.9835 | 267.4877 | 263.5611 | 265.7853 | 270.3393 | 284.1563 | 270.2441 | 259.6661 | 266.7189 | 268.6874 | 278.8636 | 268.3194 | 260.642 | 265.626 | 272.4537 | 286.5337 | 270.4901 | 261.5732 | 266.4515
268.2882 | 271.1912 | 267.6046 | 265.9394 | 266.6626 | 270.3973 | 281.2603 | 270.3623 | 262.0093 | 267.5992 | 268.7451 | 276.0216 | 268.4368 | 262.9939 | 266.5027 | 272.5122 | 283.6135 | 270.6084 | 263.9335 | 267.3309
268.52 | 272.2568 | 266.7835 | 265.4906 | 265.3491 | 270.631 | 282.3655 | 269.5327 | 261.5671 | 266.281 | 268.9773 | 277.1062 | 267.6131 | 262.5501 | 265.1899 | 272.7476 | 284.7279 | 269.778 | 263.4881 | 266.0141
269.6172 | 274.3553 | 267.4179 | 265.7708 | 265.6015 | 271.7368 | 284.5419 | 270.1737 | 261.8432 | 266.5344 | 270.0764 | 279.2421 | 268.2495 | 262.8272 | 265.4422 | 273.8621 | 286.9225 | 270.4196 | 263.7662 | 266.2672
269.0071 | 274.3917 | 267.7392 | 264.6329 | 265.9975 | 271.1218 | 284.5797 | 270.4983 | 260.722 | 266.9318 | 269.4652 | 279.2791 | 268.5718 | 261.7018 | 265.838 | 273.2424 | 286.9606 | 270.7445 | 262.6368 | 266.6642
269.3157 | 275.5332 | 268.0405 | 264.262 | 266.0796 | 271.4329 | 285.7635 | 270.8027 | 260.3566 | 267.0142 | 269.7744 | 280.4409 | 268.874 | 261.335 | 265.9201 | 273.5559 | 288.1543 | 271.0491 | 262.2688 | 266.7465
269.4089 | 276.0009 | 269.2848 | 264.0449 | 267.5284 | 271.5268 | 286.2486 | 272.0598 | 260.1428 | 268.4681 | 269.8677 | 280.9169 | 270.1222 | 261.1203 | 267.368 | 273.6505 | 288.6434 | 272.3074 | 262.0533 | 268.1989
269.7904 | 282.811 | 269.3046 | 259.1878 | 265.7749 | 271.9113 | 293.3115 | 272.0798 | 255.3575 | 266.7084 | 270.2499 | 287.8483 | 270.142 | 256.3171 | 265.6155 | 274.038 | 295.7655 | 272.3274 | 257.2329 | 266.441
269.9167 | 281.7441 | 269.4949 | 260.2859 | 266.2936 | 272.0386 | 292.205 | 272.272 | 256.4393 | 267.2289 | 270.3764 | 286.7624 | 270.3329 | 257.403 | 266.134 | 274.1663 | 294.6498 | 272.5198 | 258.3227 | 266.961
269.7928 | 284.2381 | 269.4965 | 258.0285 | 265.6444 | 271.9138 | 294.7917 | 272.2737 | 254.2153 | 266.5775 | 270.2523 | 289.3009 | 270.3345 | 255.1706 | 265.4851 | 274.0405 | 297.258 | 272.5215 | 256.0823 | 266.3102
269.4225 | 282.5092 | 270.0582 | 258.7648 | 266.8046 | 271.5405 | 292.9986 | 272.8412 | 254.9407 | 267.7417 | 269.8813 | 287.5412 | 270.898 | 255.8987 | 266.6446 | 273.6643 | 295.4499 | 273.0895 | 256.813 | 267.4732
271.3621 | 284.4547 | 270.0705 | 260.6994 | 266.312 | 273.4954 | 295.0162 | 272.8536 | 256.8467 | 267.2474 | 271.8242 | 289.5213 | 270.9103 | 257.8119 | 266.1523 | 275.6345 | 297.4845 | 273.1019 | 258.733 | 266.9795
271.4035 | 292.2632 | 271.0509 | 254.4116 | 265.5092 | 273.537 | 303.1147 | 273.8441 | 250.6518 | 266.4417 | 271.8656 | 297.4689 | 271.8937 | 251.5937 | 265.3499 | 275.6764 | 305.6507 | 274.0933 | 252.4926 | 266.1746
272.1013 | 289.6719 | 270.7188 | 257.7892 | 265.7673 | 274.2404 | 300.4272 | 273.5086 | 253.9795 | 266.7008 | 272.5647 | 294.8314 | 271.5607 | 254.9339 | 265.6079 | 276.3853 | 302.9407 | 273.7576 | 255.8448 | 266.4334
271.7212 | 290.9099 | 270.9919 | 256.0904 | 265.7884 | 273.8573 | 301.7112 | 273.7845 | 252.3058 | 266.7219 | 272.1839 | 296.0915 | 271.8345 | 253.254 | 265.629 | 275.9992 | 304.2354 | 274.0336 | 254.1588 | 266.4545
271.1276 | 289.8 | 270.718 | 255.9172 | 265.733 | 273.259 | 300.5601 | 273.5078 | 252.1351 | 266.6663 | 271.5893 | 294.9619 | 271.5598 | 253.0826 | 265.5736 | 275.3963 | 303.0747 | 273.7567 | 253.9869 | 266.399
271.2295 | 288.4066 | 270.2843 | 257.2372 | 265.5499 | 273.3617 | 299.1149 | 273.0696 | 253.4357 | 266.4826 | 271.6914 | 293.5436 | 271.1248 | 254.3881 | 265.3907 | 275.4997 | 301.6175 | 273.3181 | 255.297 | 266.2155
271.1434 | 286.7336 | 270.8904 | 258.444 | 266.7514 | 273.275 | 297.3798 | 273.682 | 254.6246 | 267.6883 | 271.6052 | 291.8408 | 271.7328 | 255.5815 | 266.5915 | 275.4123 | 299.8678 | 273.9311 | 256.4946 | 267.42
271.7981 | 289.1334 | 270.9543 | 257.6765 | 266.2053 | 273.9348 | 299.8687 | 273.7465 | 253.8685 | 267.1403 | 272.261 | 294.2834 | 271.7969 | 254.8225 | 266.0457 | 276.0773 | 302.3776 | 273.9957 | 255.733 | 266.8725
274.1596 | 289.1504 | 271.28 | 261.9482 | 266.6117 | 276.3149 | 299.8863 | 274.0756 | 258.077 | 267.5481 | 274.6265 | 294.3007 | 272.1236 | 259.0469 | 266.4518 | 278.476 | 302.3953 | 274.325 | 259.9724 | 267.2799
272.6142 | 293.3654 | 271.468 | 255.7088 | 265.748 | 274.7573 | 304.2578 | 274.2655 | 251.9299 | 266.6814 | 273.0784 | 298.5907 | 272.3121 | 252.8766 | 265.5886 | 276.9062 | 306.8034 | 274.5151 | 253.7801 | 266.414
272.7782 | 292.9679 | 271.5187 | 256.3302 | 265.9159 | 274.9226 | 303.8455 | 274.3168 | 252.5421 | 266.8498 | 273.2427 | 298.1861 | 272.3631 | 253.4911 | 265.7564 | 277.0728 | 306.3877 | 274.5664 | 254.3968 | 266.5823
272.7814 | 292.3467 | 271.851 | 256.842 | 266.4972 | 274.9258 | 303.2013 | 274.6524 | 253.0463 | 267.4332 | 273.2459 | 297.5539 | 272.6963 | 253.9972 | 266.3374 | 277.0761 | 305.738 | 274.9024 | 254.9047 | 267.1651